## Master Checklist
Subsystem | Category | Item | Qty | Priority | Status | Device / Interface | Purpose | Example Part | Product / Reference URL | Notes
Core control | Existing equipment | Main control laptop | 1 | Required | Have | System | Runs PyQt GUI, logger, device services, and possibly Agilent software |  |  | Confirm enough USB/Ethernet connectivity and power.
Core control | Existing equipment | USB dock / hub | 2 | Required | Need to Buy | System | Central edge-access connection point for USB peripherals |  |  | Prefer externally powered if many USB devices are connected.
Core control | Controller | Arduino Due | 1 | Required | Need to Buy | Pneumatics / GC sync | Main low-level controller for pneumatic valves, relays, other solenoids, and GC APG I/O | Arduino Due | https://store.arduino.cc/products/arduino-due | 7–12 V recommended external input; 3.3 V logic.
Core control | DAQ | Adafruit Feather RP2040 Adalogger | 1 | Required | Need to Buy | Analog instruments | Samples GP475 and PDR-D-1 analog outputs; two spare ADC channels remain | Adafruit PID 5980 | https://www.adafruit.com/product/5980 | Baseline uses onboard A0–A3 ADCs; ADS1115 optional.
Core control | Cable | USB data cable for Arduino Due | 1 | Required | Have / Verify | Arduino Due | Laptop ↔ Due programming/data |  |  | Match Due USB connector and dock port.
Core control | Cable | USB-C data cable for Feather | 1 | Required | Have / Verify | Feather DAQ | Laptop/dock ↔ Feather USB data + power |  |  | Must be data-capable.
Power | Main PSU | 24 V DC main power supply | 1 | Required | Need to Buy | 24 V bus | Feeds pneumatic valves, solenoids, and downstream DC/DC converters | Mean Well HDR-150-24 (example) | https://www.meanwell.com/webapp/product/search.aspx?prod=HDR-150 | Final wattage must be checked against all 24 V actuator loads.
Power | DC/DC | 24 V → 12 V regulator for Arduino Due | 1 | Required | Need to Buy | Arduino Due | Feeds Due barrel jack / VIN from main 24 V bus | Pololu D24V22F9 | https://www.pololu.com/product/2861 | Compact option; 9 V is inside Due's recommended input range.
Power | DC/DC | 24 V → 5 V DIN-rail regulator | 1 | Required | Need to Buy | Relay power | Dedicated 5 V supply for relay coils/peripherals | Mean Well DDR-15G-5 | https://www.meanwell.com/webapp/product/search.aspx?prod=DDR-15G | Keep relay coil current off Due 5 V rail.
Power | Distribution | 24 V DIN terminal blocks / distribution blocks | 1 | Required | Optional | 24 V bus | Distribute +24 V and 0 V to branch circuits |  |  | Quantity depends on number of branches.
Power | Protection | Branch fuse holders / fused terminal blocks | 1 | Required | Optional | 24 V bus | Separate fusing for pneumatics, Due converter, relay converter, other solenoids |  |  | Finalize ratings after current budget.
Power | Protection | Main DC fuse / breaker after PSU | 1 | Recommended | Optional | 24 V bus | Protects primary 24 V distribution wiring |  |  | Coordinate with PSU and conductor ampacity.
Power | Mains | Mains inlet / switch / fuse or protected power entry | 1 | Recommended | Optional | 24 V PSU | Safe AC mains connection to PSU |  |  | Use enclosed, strain-relieved, rated implementation.
Power | Grounding | DIN PE / chassis ground terminals | 1 | Recommended | Optional | System | Protective-earth bonding and organized grounding |  |  | Useful with metal DIN rail / PSU enclosure.
Power | AC distribution | Laboratory power strip / PDU for standalone instruments | 1 | Required | Have | External instruments | Powers standalone instruments and laptop as applicable |  |  | Check current rating and outlet count.
Pneumatics | Existing equipment | SMC / Orange Coast manifold with VQC1A01-51 valves | 1 | Required | Have | Pneumatics | Primary pneumatic valve manifold |  |  | 24 V valve power switched through relays/drivers.
Pneumatics | Existing equipment | 5 V relay module(s) | 1 | Required | Have | Pneumatics | Switches 24 V valve/solenoid circuits under Due control | Elegoo 4-channel 5 V relay module (existing) |  | Confirm relay channel count and JD-VCC/VCC arrangement.
Pneumatics | Actuator wiring | 24 V valve/solenoid hookup wire | 1 | Required | Have | Pneumatics | Power wiring from fused distribution through relays to valves |  |  | Gauge based on current and run length.
Pneumatics | Protection | Flyback / suppression for inductive loads | 1 | Recommended | Have | Pneumatics | Reduces solenoid switching transients |  |  | Depends on coil built-in suppression and relay topology.
Pneumatics | Other actuators | Other 24 V solenoids / crusher actuators | 1 | Required | Have | Pneumatics | Additional controlled actuators |  |  | Record each coil's rated current before finalizing PSU/fuses.
GRAPHIX | Existing instrument | Leybold GRAPHIX gauge controller | 1 | Required | Have | RS-232 | Digital pressure measurement |  | https://www.leybold.com/en-us/products/vacuum-measurement-and-control/vacuum-gauge-and-pump-controllers/graphix | Confirm exact model and baud rate.
GRAPHIX | Interface | USB → true RS-232 adapter | 1 | Required | Need to Buy | GRAPHIX | Laptop USB ↔ GRAPHIX RS-232 | USBGear USBG-232 (FTDI) | https://www.usbgear.com/usbg-232.html | Do not substitute TTL UART.
GRAPHIX | Cable | Straight-through DB9 male↔female serial cable | 1 | Required | Need to Buy | GRAPHIX | Connects RS-232 adapter to GRAPHIX rear DB9 | StarTech MXT1001MBK or equivalent | https://www.startech.com/en-us/cables/mxt1001mbk | Must be straight-through, not null-modem.
GP475 | Existing instrument | Granville-Phillips / MKS 475 Convectron controller | 1 | Required | Have | Analog | Pressure measurement |  | https://www.mks.com/f/475-convectron-vacuum-gauge-controller | 475001-00-T uses analog output.
GP475 | Connector | 3.5 mm / 1⁄8 in analog-output breakout | 1 | Required | Need to Buy | GP475 | Breaks rear analog phone jack out to field wiring | 3.5 mm plug-to-terminal adapter | https://www.adafruit.com/product/2790 | Verify exact jack contact assignment against manual.
GP475 | Cable | Shielded twisted-pair analog cable | 1 | Required | Need to Buy | GP475 | Analog signal + return to DAQ | Belden 9451 or equivalent | https://www.digikey.com/en/products/detail/belden-inc/9451-008500/28740168 | Route away from solenoid wiring.
GP475 | Analog front end | 30.1 kΩ 0.1% divider resistor | 1 | Required | Need to Buy | DAQ CH0 | Top resistor for input divider |  |  | One per populated analog channel.
GP475 | Analog front end | 10.0 kΩ 0.1% divider resistor | 1 | Required | Need to Buy | DAQ CH0 | Bottom resistor for input divider |  |  | One per populated analog channel.
GP475 | Analog front end | 100 nF capacitor | 1 | Required | Need to Buy | DAQ CH0 | Input low-pass filtering |  |  | One per populated analog channel.
PDR-D-1 | Existing instrument | MKS PDR-D-1 readout / Baratron power supply | 1 | Required | Have | Analog | Pressure measurement |  | https://clara.nz/docs/hardware-unsorted/mks/Pdr-d-man.pdf | DC OUT referenced to S GND.
PDR-D-1 | Existing instrument | Attached MKS Baratron transducer | 1 | Required | Have | PDR-D-1 | Actual pressure sensor |  |  | Record exact model and full-scale pressure.
PDR-D-1 | Cable | Shielded twisted-pair analog cable | 1 | Required | Need to Buy | PDR-D-1 | DC OUT + S GND to DAQ | Belden 9451 or equivalent | https://www.digikey.com/en/products/detail/belden-inc/9451-008500/28740168 | Order actual run length.
PDR-D-1 | Analog front end | 30.1 kΩ 0.1% divider resistor | 1 | Required | Need to Buy | DAQ CH1 | Top resistor for input divider |  |  | One per channel.
PDR-D-1 | Analog front end | 10.0 kΩ 0.1% divider resistor | 1 | Required | Need to Buy | DAQ CH1 | Bottom resistor for ADC divider |  |  | One per channel.
PDR-D-1 | Analog front end | 100 nF capacitor | 1 | Required | Need to Buy | DAQ CH1 | Input low-pass filtering |  |  | One per channel.
LAUDA | Existing instrument | LAUDA RP 290 E NTRL | 1 | Required | Have | Ethernet | Temperature control, monitoring, pump control, status/alarm access |  | https://www.lauda.de/us/products/constant-temperature-equipment/thermostats/circulation-and-process-thermostats/pro | Serial S230000611; use Ethernet process interface.
LAUDA | Network | Ethernet patch cable | 1 | Required | Have / Verify | LAUDA | LAUDA ↔ switch/dock/laptop |  |  | Cat5e/Cat6 is sufficient.
Agilent 7890A | Existing instrument | Agilent Technologies 7890A GC | 1 | Required | Have | Ethernet + APG Remote | Gas chromatography; vendor software remains primary controller |  | https://www.agilent.com/ | Keep authoritative chromatogram acquisition in Agilent software.
Agilent 7890A | Network | Ethernet patch cable | 1 | Required | Have / Verify | 7890A | GC ↔ network for OpenLab/ChemStation |  |  | Cat5e/Cat6.
Agilent 7890A | Sync interface | DB9 cable / breakout for APG REMOTE | 1 | Required | Have / Verify | 7890A ↔ Due | Carries READY/START/STOP/PREPARE synchronization |  |  | Wire only documented APG pins.
Agilent 7890A | Sync interface | Protected / optoisolated digital interface board | 1 | Recommended | TBD | 7890A ↔ Due | Buffers/protects APG signals between GC and Due |  |  | Final circuit/part selection still needs design.
Agilent 7890A | Optional analog | Shielded analog cable for SIG1/SIG2 | 1 | Optional | Optional | 7890A → Feather | Live approximate detector trace in main GUI | Belden 9451 or equivalent | https://www.digikey.com/en/products/detail/belden-inc/9451-008500/28740168 | Not a replacement for native Agilent acquisition.
Network | Infrastructure | Small Ethernet switch | 1 | Conditional | Verify | LAUDA + 7890A + laptop | Provides enough wired network ports if dock/router does not |  |  | Needed only if existing network cannot accommodate both instruments.
Network | Infrastructure | Additional Ethernet patch cable to laptop/switch uplink | 1 | Conditional | Verify | Network | Completes network topology |  |  | Depends on final topology.
Mechanical | Structure | Rear mounting plate (acrylic or polycarbonate) | 1 | Required | Need to Buy | Assembly | Structural mounting surface |  |  | Polycarbonate is more tolerant of drilled/bolted mounting.
Mechanical | Structure | Front/top clear plate | 1 | Required | Need to Buy | Assembly | Protective cover / second plate |  |  | Material and thickness TBD.
Mechanical | Structure | Plate-to-plate standoffs/spacers | 1 | Required | Need to Buy | Assembly | Separates two sheets |  |  | Length depends on tallest component.
Mechanical | PCB mounting | M2.5/M3 PCB standoff assortment | 1 | Required | Need to Buy | Due / Feather / small boards | Mounts electronics above rear plate |  |  | Match board hole sizes.
Mechanical | General mounting | M3 machine bolts + washers + nyloc nuts | 1 | Required | Need to Buy | Small brackets / cable hardware | General through-bolted mounting |  |  | Use washers against plate.
Mechanical | Heavy mounting | M4 machine bolts + washers + nyloc nuts | 1 | Required | Need to Buy | DIN rail / cable bases / PSU | Structural through-bolted mounting |  |  | Good general standard for rear plate.
Mechanical | DIN rail | 35 mm DIN rail | 1 | Recommended | Need to Buy | Power/distribution | Mounts PSU, converters, fuses, terminals |  |  | Length based on final layout.
Cable management | Permanent routing | M4 bolt-mount cable-tie bases | 10 | Required | Optional | Internal cables | Fixes cable ties to rear plate | Panduit TM2S8-C | https://www.digikey.com/en/products/detail/panduit-corp/TM2S8-C/1306625 | Use M4 through-bolts, washer and nyloc nut.
Cable management | Reusable strain relief | Reusable ratchet cable clamp near USB dock | 1 | Recommended | Optional | USB dock | Protects dock-connected cable bundle from yanks | HellermannTyton RCA90SM6 / 151-01643 | https://www.hellermanntyton.com/products/clips-clamps-and-plugs/rca90sm6/151-01643 | Fits ~6.2–13.7 mm bundle; M6 mounting.
Cable management | Routing | Slotted wiring duct | 1 | Recommended | Optional | Assembly | Organizes longer cable runs |  |  | Prefer separate signal/analog and power ducts.
Cable management | Ties | Cable ties | 1 | Required | Have | Internal cables | Works with TM2S8-C bases |  |  | Choose widths compatible with bases.
Cable management | Reusable ties | Hook-and-loop cable straps | 1 | Recommended | Optional | Service loops | Reusable bundle restraint |  |  | Good near serviceable connections.
Cable management | Termination | Wire ferrules + ferrule crimper | 1 | Recommended | Optional | DIN terminal wiring | Reliable stranded-wire terminations |  |  | Useful for 24 V distribution and converters.
Cable management | Labeling | Wire / cable labels | 1 | Recommended | Have | All wiring | Identifies both ends of removable cables |  |  | Strongly recommended.
Optional upgrades | ADC | ADS1115 16-bit ADC breakout | 1 | Optional | Optional | Analog DAQ | Higher-resolution/differential alternative to Feather ADC | Adafruit PID 1085 | https://www.adafruit.com/product/1085 | Not required for baseline pressure acquisition.
Optional upgrades | Local logging | microSD card | 1 | Optional | Optional | Feather Adalogger | Local backup data logging |  |  | Feather already has microSD socket.

## Summary
System BOM Checklist — Purchase & Verification Summary
Metric | Count
Total line items | 60
Need to Buy | 24
Verify / Have-Verify | 7
TBD | 1
Optional | 13
Critical items to resolve before final purchase
1 | 24 V PSU sizing | Record current draw of every 24 V solenoid/actuator before locking PSU and fuse ratings
2 | PDR-D-1 conversion | Record exact Baratron model and full-scale pressure
3 | Agilent APG interface | Finalize buffer/opto-isolation circuit and connector pinout before ordering interface parts

## Interface Map
Device | System Connection | Physical Interface | Power Source | GUI Integration | Key BOM Dependency
Arduino Due | Laptop + relays + GC APG | USB + GPIO | 24→12 V DC/DC | Direct USB serial | USB cable, 9 V converter, relay wiring
Feather RP2040 | Laptop + GP475 + PDR-D-1 | USB + analog A0/A1 | USB from laptop/dock | Direct USB serial | Feather, dividers, shielded analog cable
Leybold GRAPHIX | Laptop | RS-232 via USB adapter | Instrument mains | Direct serial driver | True RS-232 adapter + straight DB9 cable
GP475 | Feather A0 | 0–7 V analog | Instrument mains | Via Feather DAQ | 3.5 mm breakout + shielded pair + divider
PDR-D-1 | Feather A1 | 0–10 V analog | Instrument mains | Via Feather DAQ | Shielded pair + divider + Baratron range
LAUDA RP 290 E | Laptop/network | Ethernet | Instrument mains | Direct Ethernet ASCII commands | Ethernet cable / switch if needed
Agilent 7890A | OpenLab + Due | Ethernet + APG Remote | Instrument mains | Vendor software + sync I/O | Ethernet + APG breakout/interface
Pneumatic manifold | Arduino Due via relays | 24 V switched | 24 V main bus | Direct control via Due | 24 V PSU, relays, branch fuses